## TEST ARTICLE LOG & TEST RESULTS
DUT Part Number | DUT Serial Number | ER Number or Lot Number | Battery Vendor | Dynamic Battery Voltage at End of Session (V) | Test Result (PASS / FAIL/TML) | Comment
MT-26023-72 | 491604106171 | ER99-001 | Maxell | 2.58 | PASS | N/A
MT-26023-72 | 507255261375 | ER99-001 | Maxell | 2.457 | PASS | N/A
MT-26023-72 | 522991864908 | ER99-001 | Maxell | 2.471 | PASS | N/A
MT-26023-72 | 556672265792 | ER99-001 | Maxell | 2.558 | PASS | N/A
MT-26023-72 | 585692515774 | ER99-001 | Maxell | 2.627 | PASS | N/A
MT-26023-72 | 437934279756 | ER99-001 | Maxell | 2.496 | PASS | N/A
MT-26023-72 | 656299654840 | ER99-001 | Maxell | 2.484 | PASS | N/A
MT-26023-72 | 440057886898 | ER99-001 | Maxell | 2.534 | PASS | N/A
MT-26023-72 | 533405604165 | ER99-001 | Maxell | 2.583 | PASS | N/A
MT-26023-72 | 439444335283 | ER99-001 | Maxell | 2.585 | PASS | N/A
MT-26023-72 | 658001680974 | ER99-001 | Maxell | 2.436 | PASS | N/A
MT-26023-72 | 442414263368 | ER99-001 | Maxell | 2.533 | PASS | N/A
MT-26023-72 | 434731696050 | ER99-001 | Maxell | 2.449 | PASS | N/A
MT-26023-72 | 657824693951 | ER99-001 | Maxell | 2.4 | PASS | N/A
MT-26023-72 | 472729665856 | ER99-001 | Maxell | 2.498 | PASS | N/A
MT-26023-72 | 463938163136 | ER99-001 | Maxell | 2.626 | PASS | N/A
MT-26023-72 | 458685651898 | ER99-001 | Maxell | 2.566 | PASS | N/A
MT-26023-72 | 516598301119 | ER99-001 | Maxell | 2.488 | PASS | N/A
MT-26023-72 | 510769881521 | ER99-001 | Maxell | 2.539 | PASS | N/A
MT-26023-72 | 446899923894 | ER99-001 | Maxell | 2.512 | PASS | N/A
MT-26023-72 | 472219765172 | ER99-001 | Maxell | 2.522 | PASS | N/A
MT-26023-72 | 567349728581 | ER99-001 | Maxell | 2.629 | PASS | N/A
MT-26023-72 | 649177899591 | ER99-001 | Maxell | 2.385 | PASS | N/A
MT-26023-72 | 585703844031 | ER99-001 | Maxell | 2.554 | PASS | N/A
MT-26023-72 | 507674501708 | ER99-001 | Maxell | 2.431 | PASS | N/A
MT-26023-72 | 442083047488 | ER99-001 | Maxell | 2.612 | PASS | N/A
MT-26023-72 | 435591902143 | ER99-001 | Maxell | 2.616 | PASS | N/A
MT-26023-72 | 446931061569 | ER99-001 | Maxell | 2.482 | PASS | N/A
MT-26023-72 | 462973509702 | ER99-001 | Maxell | 2.371 | PASS | N/A
MT-26023-72 | 629445668933 | ER99-001 | Maxell | 2.315 | PASS | N/A
MT-26023-72 | 442684929717 | ER99-001 | Maxell | 2.484 | PASS | N/A
MT-26023-72 | 652706404785 | ER99-001 | Maxell | 2.44 | PASS | N/A
MT-26023-72 | 469537991346 | ER99-001 | Maxell | 2.545 | PASS | N/A
MT-26023-72 | 454299321934 | ER99-001 | Maxell | 2.564 | PASS | N/A
MT-26023-72 | 472524309949 | ER99-001 | Maxell | 2.533 | PASS | N/A
MT-26023-72 | 432169948484 | ER99-001 | Maxell | 2.542 | PASS | N/A
MT-26023-72 | 470249504447 | ER99-001 | Maxell | 2.49 | PASS | N/A
MT-26023-72 | 521924382389 | ER99-001 | Maxell | 2.634 | PASS | N/A
MT-26023-72 | 579117415868 | ER99-001 | Maxell | 2.508 | PASS | N/A
MT-26023-72 | 574864854964 | ER99-001 | Maxell | 2.569 | PASS | N/A
MT-26023-72 | 584035204928 | ER99-001 | Maxell | 2.615 | PASS | N/A
MT-26023-72 | 499889514574 | ER99-001 | Maxell | 2.585 | PASS | N/A
MT-26023-72 | 553877515080 | ER99-001 | Maxell | 2.483 | PASS | N/A
MT-26023-72 | 439189719872 | ER99-001 | Maxell | 2.407 | PASS | N/A
MT-26023-72 | 555101230527 | ER99-001 | Maxell | 2.363 | PASS | N/A
MT-26023-72 | 906341695928 | ER99-001 | Maxell | 2.436 | PASS | N/A
MT-26023-72 | 448990036300 | ER99-001 | Maxell | 2.42 | PASS | N/A
MT-26023-72 | 476589491009 | ER99-001 | Maxell | 2.544 | PASS | N/A
MT-26023-72 | 637899162183 | ER99-001 | Maxell | 2.624 | PASS | N/A
MT-26023-72 | 657014909505 | ER99-001 | Maxell | 2.488 | PASS | N/A
MT-26023-72 | 582216277440 | ER99-001 | Maxell | 2.433 | FAIL | N/A
MT-26023-72 | 567975907656 | ER99-001 | Maxell | 2.456 | PASS | N/A
MT-26023-72 | 489357899075 | ER99-001 | Maxell | 2.376 | PASS | N/A
MT-26023-72 | 472045017276 | ER99-001 | Maxell | 2.536 | PASS | N/A
MT-26023-72 | 535931921590 | ER99-001 | Maxell | 2.352 | PASS | N/A
MT-26023-72 | 608790621887 | ER99-001 | Maxell | 2.382 | PASS | N/A
MT-26023-72 | 655976628657 | ER99-001 | Maxell | 2.426 | PASS | N/A
MT-26023-72 | 437088526782 | ER99-001 | Maxell | 2.491 | PASS | N/A
MT-26023-72 | 437506204339 | ER99-001 | Maxell | 2.507 | PASS | N/A
MT-26023-72 | 440105761605 | ER99-001 | Maxell | 2.474 | PASS | N/A
MT-26023-72 | 434189566026 | ER99-001 | Maxell | 2.552 | PASS | N/A
MT-26023-72 | 579620984644 | ER99-001 | Maxell | 2.447 | PASS | N/A
MT-26023-72 | 500700248761 | ER99-001 | Maxell | 2.526 | PASS | N/A
MT-26023-72 | 648596081083 | ER99-001 | Maxell | 2.548 | PASS | N/A
MT-26023-72 | 639568109894 | ER99-001 | Maxell | 2.407 | PASS | N/A
MT-26023-72 | 437305275319 | ER99-001 | Maxell | 2.465 | PASS | N/A
MT-26023-72 | 479997855044 | ER99-001 | Maxell | 2.384 | PASS | N/A
MT-26023-72 | 582293061955 | ER99-001 | Maxell | 2.655 | PASS | N/A
MT-26023-72 | 586269423025 | ER99-001 | Maxell | 2.579 | PASS | N/A
MT-26023-72 | 571540396216 | ER99-001 | Maxell | 2.563 | PASS | N/A
MT-26023-72 | 510789560133 | ER99-001 | Maxell | 2.516 | PASS | N/A
MT-26023-72 | 459104917686 | ER99-001 | Maxell | 2.51 | PASS | N/A
MT-26023-72 | 651111897934 | ER99-001 | Maxell | 2.512 | PASS | N/A
MT-26023-72 | 444240564300 | ER99-001 | Maxell | 2.597 | PASS | N/A
MT-26023-72 | 435519728207 | ER99-001 | Maxell | 2.583 | PASS | N/A
MT-26023-72 | 555267920206 | ER99-001 | Maxell | 2.324 | PASS | N/A
MT-26023-72 | 650761263292 | ER99-001 | Maxell | 2.497 | PASS | N/A
MT-26023-72 | 660712534207 | ER99-001 | Maxell | 2.514 | PASS | N/A
MT-26023-72 | 443582721211 | ER99-001 | Maxell | 2.642 | PASS | N/A
MT-26023-72 | 586452511413 | ER99-001 | Maxell | 2.5 | PASS | N/A
MT-26023-72 | 519734523460 | ER99-001 | Maxell | 2.496 | PASS | N/A
MT-26023-72 | 438658141255 | ER99-001 | Maxell | 2.475 | PASS | N/A
MT-26023-72 | 652267016271 | ER99-001 | Maxell | 2.442 | PASS | N/A
MT-26023-72 | 910955823688 | ER99-001 | Maxell | 2.431 | PASS | N/A
MT-26023-72 | 515813159347 | ER99-001 | Maxell | 2.493 | PASS | N/A
MT-26023-72 | 441675500471 | ER99-001 | Maxell | 2.507 | PASS | N/A
MT-26023-72 | 912514997829 | ER99-001 | Maxell | 2.437 | PASS | N/A
MT-26023-72 | 586114883908 | ER99-001 | Maxell | 2.567 | PASS | N/A
MT-26023-72 | 585533693376 | ER99-001 | Maxell | 2.587 | PASS | N/A
MT-26023-72 | 521038352965 | ER99-001 | Maxell | 2.542 | PASS | N/A
MT-26023-72 | 442718897729 | ER99-001 | Maxell | 2.493 | PASS | N/A
MT-26023-72 | 582876976063 | ER99-001 | Maxell | 2.488 | PASS | N/A
MT-26023-72 | 497577220930 | ER99-001 | Maxell | 2.476 | PASS | N/A
MT-26023-72 | 472732732486 | ER99-001 | Maxell | 2.526 | PASS | N/A
MT-26023-72 | 484152253507 | ER99-001 | Maxell | 2.478 | PASS | N/A
MT-26023-72 | 652737148228 | ER99-001 | Maxell | 2.583 | PASS | N/A
MT-26023-72 | 472578431931 | ER99-001 | Maxell | 2.512 | PASS | N/A
MT-26023-72 | 514531185334 | ER99-001 | Maxell | 2.571 | PASS | N/A
MT-26023-72 | 437908031933 | ER99-001 | Maxell | 2.428 | PASS | N/A
MT-26023-72 | 443951240628 | ER99-001 | Maxell | 2.493 | PASS | N/A
MT-26023-72 | 498635995826 | ER99-001 | Maxell | 2.436 | PASS | N/A
MT-26023-72 | 589796483262 | ER99-001 | Maxell | 2.574 | PASS | N/A
MT-26023-72 | 532620019785 | ER99-001 | Maxell | 2.591 | PASS | N/A
MT-26023-72 | 657838644043 | ER99-001 | Maxell | 2.623 | PASS | N/A
MT-26023-72 | 584032681659 | ER99-001 | Maxell | 2.534 | PASS | N/A
MT-26023-72 | 632975197513 | ER99-001 | Maxell | 2.419 | PASS | N/A
MT-26023-72 | 639770451643 | ER99-001 | Maxell | 2.591 | PASS | N/A
MT-26023-72 | 641073099957 | ER99-001 | Maxell | 2.525 | PASS | N/A
MT-26023-72 | 512557031355 | ER99-001 | Maxell | 2.612 | PASS | N/A
MT-26023-72 | 651441806924 | ER99-001 | Maxell | 2.252 | PASS | N/A
MT-26023-72 | 659359830717 | ER99-001 | Maxell | 2.48 | PASS | N/A
MT-26023-72 | 475655186243 | ER99-001 | Maxell | 2.6 | PASS | N/A
MT-26023-72 | 437725299641 | ER99-001 | Maxell | 2.425 | PASS | N/A
MT-26023-72 | 511917828029 | ER99-001 | Maxell | 2.492 | PASS | N/A
MT-26023-72 | 523499826108 | ER99-001 | Maxell | 2.595 | PASS | N/A
MT-26023-72 | 657309003195 | ER99-001 | Maxell | 2.478 | PASS | N/A
MT-26023-72 | 510254373967 | ER99-001 | Maxell | 2.613 | PASS | N/A
MT-26023-72 | 505720883643 | ER99-001 | Maxell | 2.447 | PASS | N/A
MT-26023-72 | 913239727690 | ER99-001 | Maxell | 2.628 | PASS | N/A
MT-26023-72 | 507540360376 | ER99-001 | Maxell | 2.524 | PASS | N/A
MT-26023-72 | 479010884023 | ER99-002 | Panasonic | 2.495 | PASS | N/A
MT-26023-72 | 435514064320 | ER99-002 | Panasonic | 2.307 | PASS | N/A
MT-26023-72 | 483884698550 | ER99-002 | Panasonic | 2.461 | PASS | N/A
MT-26023-72 | 450154472520 | ER99-002 | Panasonic | 2.464 | PASS | N/A
MT-26023-72 | 523355021130 | ER99-002 | Panasonic | 2.516 | PASS | N/A
MT-26023-72 | 434959724108 | ER99-002 | Panasonic | 2.497 | PASS | N/A
MT-26023-72 | 517831688380 | ER99-002 | Panasonic | 2.595 | PASS | N/A
MT-26023-72 | 912338829634 | ER99-002 | Panasonic | 2.545 | PASS | N/A
MT-26023-72 | 579923182268 | ER99-002 | Panasonic | 2.429 | PASS | N/A
MT-26023-72 | 512404453301 | ER99-002 | Panasonic | 2.407 | PASS | N/A
MT-26023-72 | 484557559222 | ER99-002 | Panasonic | 2.629 | PASS | N/A
MT-26023-72 | 910044715454 | ER99-002 | Panasonic | 2.421 | PASS | N/A
MT-26023-72 | 516117188936 | ER99-002 | Panasonic | 2.452 | PASS | N/A
MT-26023-72 | 650923673531 | ER99-002 | Panasonic | 2.499 | PASS | N/A
MT-26023-72 | 909121264829 | ER99-002 | Panasonic | 2.393 | PASS | N/A
MT-26023-72 | 485674393272 | ER99-002 | Panasonic | 2.273 | PASS | N/A
MT-26023-72 | 647858908229 | ER99-002 | Panasonic | 2.361 | PASS | N/A
MT-26023-72 | 466656097465 | ER99-002 | Panasonic | 2.463 | PASS | N/A
MT-26023-72 | 518105105851 | ER99-002 | Panasonic | 2.569 | PASS | N/A
MT-26023-72 | 463261914956 | ER99-002 | Panasonic | 2.385 | PASS | N/A
MT-26023-72 | 910856974156 | ER99-002 | Panasonic | 2.448 | PASS | N/A
MT-26023-72 | 656606594491 | ER99-002 | Panasonic | 2.456 | PASS | N/A
MT-26023-72 | 469966070969 | ER99-002 | Panasonic | 2.416 | PASS | N/A
MT-26023-72 | 570281018183 | ER99-002 | Panasonic | 2.54 | FAIL | N/A
MT-26023-72 | 526321745079 | ER99-002 | Panasonic | 2.519 | PASS | N/A
MT-26023-72 | 459424308913 | ER99-002 | Panasonic | 2.62 | PASS | N/A
MT-26023-72 | 507971900234 | ER99-002 | Panasonic | 2.427 | PASS | N/A
MT-26023-72 | 437496856138 | ER99-002 | Panasonic | 2.506 | PASS | N/A
MT-26023-72 | 644490839743 | ER99-002 | Panasonic | 2.583 | PASS | N/A
MT-26023-72 | 499543941451 | ER99-002 | Panasonic | 2.488 | PASS | N/A
MT-26023-72 | 912062013624 | ER99-002 | Panasonic | 2.502 | PASS | N/A
MT-26023-72 | 581982952114 | ER99-002 | Panasonic | 2.483 | PASS | N/A
MT-26023-72 | 450749117876 | ER99-002 | Panasonic | 2.613 | PASS | N/A
MT-26023-72 | 653761205181 | ER99-002 | Panasonic | 2.592 | PASS | N/A
MT-26023-72 | 471781984184 | ER99-002 | Panasonic | 2.465 | PASS | N/A
MT-26023-72 | 585132294208 | ER99-002 | Panasonic | 2.397 | PASS | N/A
MT-26023-72 | 656648158903 | ER99-002 | Panasonic | 2.616 | PASS | N/A
MT-26023-72 | 654678590141 | ER99-002 | Panasonic | 2.461 | PASS | N/A
MT-26023-72 | 654970007991 | ER99-002 | Panasonic | 2.553 | PASS | N/A
MT-26023-72 | 570209750201 | ER99-002 | Panasonic | 2.531 | PASS | N/A
MT-26023-72 | 909139976001 | ER99-002 | Panasonic | 2.511 | PASS | N/A
MT-26023-72 | 911954169523 | ER99-002 | Panasonic | 2.423 | PASS | N/A
MT-26023-72 | 470740144716 | ER99-002 | Panasonic | 2.481 | PASS | N/A
MT-26023-72 | 909766082236 | ER99-002 | Panasonic | 2.491 | PASS | N/A
MT-26023-72 | 443515446854 | ER99-002 | Panasonic | 2.538 | PASS | N/A
MT-26023-72 | 470151714749 | ER99-002 | Panasonic | 2.582 | PASS | N/A
MT-26023-72 | 462284832699 | ER99-002 | Panasonic | 2.515 | PASS | N/A
MT-26023-72 | 502805865664 | ER99-002 | Panasonic | 2.559 | PASS | N/A
MT-26023-72 | 519061555643 | ER99-002 | Panasonic | 2.394 | PASS | N/A
MT-26023-72 | 557606587581 | ER99-002 | Panasonic | 2.51 | PASS | N/A
MT-26023-72 | 461947184457 | ER99-002 | Panasonic | 2.58 | PASS | N/A
MT-26023-72 | 523326185280 | ER99-002 | Panasonic | 2.484 | PASS | N/A
MT-26023-72 | 433673915827 | ER99-002 | Panasonic | 2.575 | FAIL | N/A
MT-26023-72 | 474128196800 | ER99-002 | Panasonic | 2.565 | PASS | N/A
MT-26023-72 | 436135314099 | ER99-002 | Panasonic | 2.572 | PASS | N/A
MT-26023-72 | 657835503432 | ER99-002 | Panasonic | 2.595 | PASS | N/A
MT-26023-72 | 460350611788 | ER99-002 | Panasonic | 2.525 | PASS | N/A
MT-26023-72 | 554860075593 | ER99-002 | Panasonic | 2.597 | PASS | N/A
MT-26023-72 | 506319163456 | ER99-002 | Panasonic | 2.573 | PASS | N/A
MT-26023-72 | 509792427957 | ER99-002 | Panasonic | 2.558 | PASS | N/A
MT-26023-72 | 485632222393 | ER99-002 | Panasonic | 2.646 | PASS | N/A
MT-26023-72 | 657147655091 | ER99-002 | Panasonic | 2.546 | PASS | N/A
MT-26023-72 | 457987989581 | ER99-002 | Panasonic | 2.255 | PASS | N/A
MT-26023-72 | 514405103436 | ER99-002 | Panasonic | 2.528 | PASS | N/A
MT-26023-72 | 554789626432 | ER99-002 | Panasonic | 2.414 | PASS | N/A
MT-26023-72 | 651192634545 | ER99-002 | Panasonic | 2.433 | PASS | N/A
MT-26023-72 | 658448161204 | ER99-002 | Panasonic | 2.512 | PASS | N/A
MT-26023-72 | 438006750140 | ER99-002 | Panasonic | 2.494 | PASS | N/A
MT-26023-72 | 440621489855 | ER99-002 | Panasonic | 2.569 | PASS | N/A
MT-26023-72 | 437337878458 | ER99-002 | Panasonic | 2.583 | PASS | N/A
MT-26023-72 | 500601023411 | ER99-002 | Panasonic | 2.576 | PASS | N/A
MT-26023-72 | 461603547314 | ER99-002 | Panasonic | 2.501 | PASS | N/A
MT-26023-72 | 486495389881 | ER99-002 | Panasonic | 2.529 | PASS | N/A
MT-26023-72 | 658203284045 | ER99-002 | Panasonic | 2.477 | PASS | N/A
MT-26023-72 | 646005000256 | ER99-002 | Panasonic | 2.565 | PASS | N/A
MT-26023-72 | 581334872641 | ER99-002 | Panasonic | 2.446 | PASS | N/A
MT-26023-72 | 439868997188 | ER99-002 | Panasonic | 2.569 | PASS | N/A
MT-26023-72 | 518705780035 | ER99-002 | Panasonic | 2.525 | PASS | N/A
MT-26023-72 | 526045444285 | ER99-002 | Panasonic | 2.561 | PASS | N/A
MT-26023-72 | 466390011456 | ER99-002 | Panasonic | 2.562 | PASS | N/A
MT-26023-72 | 910556955835 | ER99-002 | Panasonic | 2.515 | PASS | N/A
MT-26023-72 | 553786822976 | ER99-002 | Panasonic | 2.407 | PASS | N/A
MT-26023-72 | 554773854272 | ER99-002 | Panasonic | 2.435 | PASS | N/A
MT-26023-72 | 525914589362 | ER99-002 | Panasonic | 2.446 | PASS | N/A
MT-26023-72 | 577902835779 | ER99-002 | Panasonic | 2.412 | PASS | N/A
MT-26023-72 | 908453371833 | ER99-002 | Panasonic | 2.539 | PASS | N/A
MT-26023-72 | 432394089663 | ER99-002 | Panasonic | 2.6 | PASS | N/A
MT-26023-72 | 432362045621 | ER99-002 | Panasonic | 2.544 | PASS | N/A
MT-26023-72 | 654858751416 | ER99-002 | Panasonic | 2.447 | PASS | N/A
MT-26023-72 | 441269727816 | ER99-002 | Panasonic | 2.504 | PASS | N/A
MT-26023-72 | 437733075640 | ER99-002 | Panasonic | 2.538 | PASS | N/A
MT-26023-72 | 656929641805 | ER99-002 | Panasonic | 2.492 | PASS | N/A
MT-26023-72 | 554766974279 | ER99-002 | Panasonic | 2.603 | PASS | N/A
MT-26023-72 | 654202542403 | ER99-002 | Panasonic | 2.564 | PASS | N/A
MT-26023-72 | 608508424629 | ER99-002 | Panasonic | 2.441 | PASS | N/A
MT-26023-72 | 472551611575 | ER99-002 | Panasonic | 2.533 | PASS | N/A
MT-26023-72 | 441702553268 | ER99-002 | Panasonic | 2.586 | PASS | N/A
MT-26023-72 | 514801081011 | ER99-002 | Panasonic | 2.5 | PASS | N/A
MT-26023-72 | 568603891779 | ER99-002 | Panasonic | 2.442 | PASS | N/A
MT-26023-72 | 472754830527 | ER99-002 | Panasonic | 2.462 | PASS | N/A
MT-26023-72 | 658222815822 | ER99-002 | Panasonic | 2.565 | PASS | N/A
MT-26023-72 | 471556789937 | ER99-002 | Panasonic | 2.584 | PASS | N/A
MT-26023-72 | 460816012101 | ER99-002 | Panasonic | 2.545 | PASS | N/A
MT-26023-72 | 440934761794 | ER99-002 | Panasonic | 2.499 | PASS | N/A
MT-26023-72 | 497497185723 | ER99-002 | Panasonic | 2.465 | PASS | N/A
MT-26023-72 | 909593946295 | ER99-002 | Panasonic | 2.335 | PASS | N/A
MT-26023-72 | 514018292804 | ER99-002 | Panasonic | 2.555 | PASS | N/A
MT-26023-72 | 518864738114 | ER99-002 | Panasonic | 2.463 | PASS | N/A
MT-26023-72 | 909466941883 | ER99-002 | Panasonic | 2.392 | PASS | N/A
MT-26023-72 | 579118122829 | ER99-002 | Panasonic | 2.431 | PASS | N/A
MT-26023-72 | 655416171079 | ER99-002 | Panasonic | 2.481 | PASS | N/A
MT-26023-72 | 647751071161 | ER99-002 | Panasonic | 2.397 | PASS | N/A
MT-26023-72 | 910394124622 | ER99-002 | Panasonic | 2.551 | PASS | N/A
MT-26023-72 | 578248138930 | ER99-002 | Panasonic | 2.398 | PASS | N/A
MT-26023-72 | 492950091849 | ER99-002 | Panasonic | 2.466 | PASS | N/A
MT-26023-72 | 472684369475 | ER99-002 | Panasonic | 2.41 | PASS | N/A
MT-26023-72 | 498337492047 | ER99-002 | Panasonic | 2.523 | PASS | N/A
MT-26023-72 | 652257513657 | ER99-002 | Panasonic | 2.572 | PASS | N/A
MT-26023-72 | 580086246581 | ER99-002 | Panasonic | 2.567 | PASS | N/A
MT-26023-72 | 645591309647 | ER99-002 | Panasonic | 2.398 | PASS | N/A
MT-26023-72 | 464617824646 | ER99-002 | Panasonic | N/A | TML | N/A

## EQUIPMENT LOG
Applicable Test Procedure or Protocol | Equipment Number | Description | Calibration Due Date | Comment
| EQ-901061-001 | Arbin BT2000 & BT2100 & HSP21044 Battery Test Systems. | 2025-02-21
| EQ-901061-002 | Arbin BT2000 & BT2100 & HSP21044 Battery Test Systems. | 2025-02-21
| EQ-901061-003 | Arbin BT2000 & BT2100 & HSP21044 Battery Test Systems. | 2025-02-21
| EQ-901061-004 | Arbin BT2000 & BT2100 & HSP21044 Battery Test Systems. | 2025-02-21
| EQ-901061-10010 | Arbin BT2000 & BT2100 & HSP21044 Battery Test Systems. | 2025-01-18
| EQ-900946-001 | Keysight N6705B DC Power Analyzer Mainframe | N/A | No calibration required. Measurement modules EQ-900947 are calibrated below.
| EQ-900947-031 | Keysight N6781A 2-Quadrant Module | 2025-07-31
| EQ-900947-032 | Keysight N6781A 2-Quadrant Module | 2025-07-31
| EQ-900947-033 | Keysight N6781A 2-Quadrant Module | 2025-07-31
| EQ-900947-035 | Keysight N6781A 2-Quadrant Module | 2025-07-31
| EQ-900523-061 | Vaisala Temperature and RH Data Recorder | 2025-01-31
| EQ-900888-007 | TestEquity Temperature/Humidity Chamber Model 123HS/ 123HS-EX | 2025-05-31
| EQ-900888-005 | TestEquity Temperature/Humidity Chamber Model 123HS/ 123HS-EX | 2025-04-30
| EQ-900888-006 | TestEquity Temperature/Humidity Chamber Model 123HS/ 123HS-EX | 2025-05-31
| EQ-900523-058 | Vaisala Temperature and RH Data Recorder | 2025-03-31
| EQ-900523-064 | Vaisala Temperature and RH Data Recorder | 2025-03-31

## SOFTWARE LOG
Applicable Test Procedure or Protocol | Software Part Number | Revision | Description

## MATERIAL LOG
Applicable Test Procedure or Protocol | Material Part Number | Revision | Description

## TEST METHOD LOSSES
# | PROTOCOL NOTE TYPE | TEST UNIT SN or ID | PROTOCOL STEP / SECTION | REPLACEMENT(S) | OBSERVATIONS
1 | TEST METHOD LOSS | 454517824545, 393917822929 | 13.1 | 949417829494, 757517827575 | Units were punctured during deployment by the operator.
2 | TEST METHOD LOSS | 464617824646 | 13.9 | N/A | Test unit were depleted for 24 hours past the maximum depleted time due to an error in calculating the depletion time. Units were not replaced as there was sufficient samples to meet the minimum sample size.

## DEVIATIONS
# | PROTOCOL NOTE TYPE | TEST UNIT SN or ID | PROTOCOL STEP / SECTION | ORIGINAL PROCEDURE STEP | DEVIATION | IMPACT TO TEST RESULTS OR REASON FOR NO IMPACT
1 | DEVIATION | All | 12.1 | After the battery predischarge, the batteries are built into G7 GSS Wearables with RC instead of a sensor, and with tacking adhesive and encapsulant epoxy. | Units did not go through the manufacturing process to get exposed to the process temperatures of the conductive epoxy. To account for the potential impact of this step on the battery life, the devices were placed into an oven at 80C / 60 min (uncontrolled RH). | The temperature and exposure time simulated represents the FASTr production line recipe RCP-1000184, which specifies a temperature of 78 C for 50 minutes. This simulated step represented the temperatures seen in manufacturing and will not impact the test results.
2 | DEVIATION | All | 11.3.2, 11.3.3 | 11.3.2 Using the Arbin battery test equipment, set the equipment to discharge the batteries for 3.47 mAh at a constant current of 0.100 mA. 11.3.3 Place batteries on the Arbin and start the test. | To account for the missing battery depletion, the average transmitter sleep current from PTL-903600 Appendix B was used to determine the missing capacity to deplete. The missing capacity to deplete was added as test session time at the end of the test. | This step adds variability to the results as the Arbin tester provides a constant capacity depletion for all batteries, while each PCBA will have a different current draw to deplete the batteries. However, the added variability was mitigated by adding 5 days to the protocol execution.

## DEFECTIVE UNITS
# | PROTOCOL NOTE TYPE | TEST UNIT SN or ID | PROTOCOL STEP / SECTION | RELATED REQUIREMENT(S) | OBSERVATIONS
1 | DEVIATION | 570281018183, 582216277440, 433673915827 | 13.5.3 | 110596, 110672 | The test article failed to communicate with the display device at 22.66 days to retreive the database from the transmitter. The Analysis is captured in Jira Ticket SHAD-2. Review of manufacturing logs confirms the issue is not related to the manufacturing of the PCBA. The review of the test method found that a database download does not represent the data download method done by display devices which is more intensive than the private data download method used by the displays. The units could not be retested due to the desctructive nature of the investigation and are considered failures for this test. However the protocol shouild be updated to reflect that cammunication with the display device should be done to retreive private data instead of the device database, as a private data download represents the way the display device request data to the transmitters.